High-level requirements for SPINS PSD calibration module

Main Goals:

- Determine the effective final energy at each horizontal pixel location for the PSD.

- Determine the counting efficiency at each horizontal pixel location for the PSD.

General notes/capabilities:

- Read a single SPINS ICP formated data file (.ng5) created with PSD running in a+ mode - this is the raw calibration data file.

- It is expected that the raw calibration file must have been collected using the ICP Q-buffer mode.

- Fit a Gaussian peak to each of the 'data points' in the dataset. Each data point consists of 256 intensity values measured at each horizontal pixel of the PSD (integration is performed along the 256 vertical pixels).

- Assign an effective final energy (energy transfer + nominal final energy) to the peak position (pixel) of the fitted Gaussian.

- Use results for all data points to determine an Ef for any pixel by interpolation.

- Calculate the integrated intensity (sum over the nos of data points) at each pixel and and correct (by dividing) for the rate of change of Ef with respect to pixel location. This provides a measure of the sensitivity of the PSD at each pixel location.

Display Capabilities:

- Display the raw data as an image using energy transfer (or effective final energy) and horizontal pixel nos as the y- and x-axis respectively. Allow zooming capability.

- Plot data + Gaussian fit for each data point. User selects group (data point) to display using a slider. Allow zooming capability.

- Plot calculated Ef vs pixel. Allow zooming capability.

- Plot calculated efficiencies vs pixel. Allow zooming capability.

Configurable parameters:

- Minimum peak intensity required for data point to be fitted with Gaussian and hence used in determining Ef.

- The pixel range to be used in the efficiency calculation. The integrated counts from pixels outside the specified range are assume to be 'background' that is subtracted off.

Outputs:

- A single ASCII file containing the following details:

  - raw data filename used

  - number of horizontal pixels in the PSD.

  - the effective final energy at each pixel

  - the counting efficiency at each pixel

- jpg/png/ps output from any of the graphic displays mentioned above.